# 4 Channel EMAT Pulser

Instruction Manual

# EW270

18 JAN 2017

# Description

The unit will drive four EMATs simultaneously. The relative delays can be adjusted via a serial link, providing timing control with 2.5ns resolution.

A SYNC OUT connector is provided on the front panel for timing reference.

A front panel switch for HT SUPPLY ON/OFF is provided, so that the unit can be set up first and then the EMATs can be switched on.

Output pulse voltage is 850V.

# Command List

*IDN?		Show unit id

?		Show current setup

HELP		Show help page

DEL 	<0-3/all>   <t>[ns/us]		set delay for specified channel(s), e.g. DEL 0 50ns

EN	<0-3/all>			enable specified channel(s), e.g. EN 3

DIS 	<0-3/all>			disable specified channel(s), e.g. DIS 0

REP	<f>				set rep rate frequency in Hz, e.g. REP 10

PW	<t> [ns/us]			set pulse width into driver circuit, e.g. PW 1us

RPW	<t> [ns/us]			set pulse width of SYNC OUTPUT pulse (front panel BNC)

STO	<0-15>				store current setup to flash memory, e.g. STO 1

RCL	<0-15>				recall settings from flash memory, e.g. RCL 2

NOTES:

- Flash memory 0 setup will be recalled at power-up

- There will be a small delay between ref pulse and output pulse due to driver circuit delays (around 100ns or so)

- Command interface is not case sensitive, except where units are ambiguous (‘m’ / ‘M’)

- Command & response termination is “\r\n” (0x0D, 0x0A)

# Specification

Output voltage		850V

Pulse capacitance	470nF

Delay resolution	2.5ns

Output Connector	N-type

Max rep rate		300Hz (software limited)

Command interface	RS232 (8 data, no parity, 1 stop, 230400 Baud, 9 way D socket)

CAUTION!!		DO NOT HANDLE EMAT CONNECTORS WITH HT SUPPLY ON

CAUTION!!		FIT BLANKING PLUGS TO UNUSED OUTPUTS